## FDA-ISO
Click here to schedule your free personal demo of Greenlight Guru's Medical Device QMS Software today.
FDA/ISO QMS Audit Checklist by Greenlight Guru
Audit #: |  |  |  | NC = Non-Conformance OFI = Opportunity for Improvement PP = Positive Practice A = Acceptable
Dates:
Lead Auditor:
Item | Subsystem / Assessment Detail | FDA / ISO reference | Auditor Notes | Auditor Observation | Objective Evidence | NC, OFI, PP, or A?
Management Controls (main subsystem)
1 | ensure Quality Manual defines scope of QMS, procedures (or reference to) within QMS, and description of the interaction of processes within QMS | ISO 13485:2016: 4.1, 4.2.2 | review quality manual; review QMS metrics; review critical processes and procedures
2 | verify criteria and methods are in place to monitor and control processes for effectiveness | ISO 13485:2016: 4.1.3(a), 4.2.1(d), 8.4 | review QMS metrics; review management reviews
3 | Verify firm has established and conducts Management Reviews, at least annually | ISO 13485:2016: 5.1(d), 5.6; 21 CFR 820.5, 820.20(c) | request procedure in advance; review management reviews
4 | confirm management reviews examine suitability and effectiveness of quality systems, improvements needed because of customer requirements, and resource needs | ISO 13485:2016: 4.1.3(c), 5.6.1, 5.6.3, 6.1, 8.4; 21 CFR 820.20(c) | review procedure
5 | ensure management review addresses audit results, customer feedback, process performance, CAPAs, previous management reviews, changes to QMS, recommendations for improvement, and new or revised regulatory requirements | ISO 13485:2016: 5.6.2 | review management reviews
6 | verify firm has established a Quality Manual and Quality System Procedures and Instructions that are appropriate | ISO 13485:2016: 4.1.2(a), 4.2.1(b), (c); 21 CFR 820.5, 820.20(c), (d), (e), 820.22 | request quality manual and procedures in advance; review documents
7 | Verify firm has established Quality Plan | ISO 13485:2016: 4.2.1(d), 5.4; 21 CFR 820.20(d) | request quality plan in advance
8 | confirm that Quality Planning addresses QMS needs and Quality Objectives | ISO 13485:2016: 5.4.2; 21 CFR 820.20(a), (d) | review quality plan
9 | verify firm has implemented Quality Policy and Quality Objectives | ISO 13485:2016: 4.2.1(a), 5.1(b), (c), 5.3, 5.4.1 21 CFR 820.20(a), (d) | interview employees about quality policy; review training records
10 | Verify firm has established Quality Audit procedures and conducts audits | ISO 13485:2016: 4.2, 8.2.4; 21 CFR 820.20(c), 820.22 | request procedure in advance; review audit schedule and documents; review auditor training
11 | ensure quality audits examine compliance and effectiveness | ISO 13485:2016: 4.1.3(c), 4.2.1(d), 8.2.4; 21 CFR 820.22 | review procedure; review audit records
12 | verify that auditors are trained | ISO 13485:2016: 6.2, 8.2.4; 21 CFR 820.22 | review audit records; review training records
13 | ensure that audits are conducted by objective parties | ISO 13485:2016: 8.2.4; 21 CFR 820.22 | review audit records; review training records
14 | confirm quality audits are linked to CAPA | ISO 13485:2016: 8.2.4; 21 CFR 820.22, 820.100 | review procedures
15 | Review organizational structure of firm; confirm resources are available to support processes | ISO 13485:2016: 4.1.3(b), 5.1(e), 5.5.1, 5.5.2, 6.1, 6.2; 21 CFR 820.20(b), 820.25 | request organizational chart(s) in advance
16 | verify firm has defined a management representative with executive responsibility for implementing and reporting quality management system | ISO 13485:2016: 5.1, 5.5.1, 5.5.2, 6.1, 6.2; 21 CFR 820.20(b)(3), 820.25 | ask management representative to identify responsibility for: -changes to procedures, device designs, manufacturing processes -review of quality audit results -oversight and interaction with CAPA activities
17 | verify appropriate responsibilities , authority, and resources are in place for quality system activities | ISO 13485:2016: 5.1(e), 5.5.1, 5.5.2, 6.1, 6.2; 21 CFR 820.5(b)(1)-(2), 820.20(b), 820.25 | interview management representative about resource allocation
18 | verify firm has established procedures for identifying training needs; ensure personnel are trained to perform assigned responsibilities | ISO 13485:2016: 6.2; 21 CFR 820.25(b) | review procedures; review training records
19 | AT AUDIT CONCLUSION . . . Determine if executive management ensures adequate and effective quality system is implemented. Ensure management is committed to and communicates importance of meeting customer requirements, regulatory requirements, and QMS. | ISO 13485:2016: 5.1(a), 5.2, 5.5.3 | interview executive management; provide confirmation or failures of quality system; review other subsystems and return to management controls
Design & Development / Design Controls (main subsystem)
1 | verify products are subject to design controls | ISO 13485:2016: 7.1, 7.3; 21 CFR 820.30(a) | review procedure; review products
2 | verify design control and risk management procedures are established and applied | ISO 13485:2016: 7.3; 21 CFR 820.30(a) - (j) | ensure procedures address all design control elements
3 | ensure design and development stages are identified; confirm that review, verification, validation, and design transfer activities at each stage are appropriate; verify responsibilities for design and development are defined | ISO 13485:2016: 7.3.2; 21 CFR 820.30 | review procedures
4 | select a design project |  | selection criteria: -contains software -single product focus -risk based -result of complaints, problems -most recent -cover product range -recent 510(k), PMA, CE mark
5 | review the project design & development plan, responsibilities, and interfaces | ISO 13485:2016: 7.3.2; 21 CFR 820.30(b) | review procedure; assess plan's -milestones -phases -responsibilities -risk management -interfaces
6 | verify design & development plan is updated, reviewed, and approved | ISO 13485:2016: 7.3.2; 21 CFR 820.30(b) | review plan revisions; review and approval procedures
7 | confirm design input requirements were established, reviewed, and approved; ensure customer requirements are captured; ensure inputs include functional, performance, safety, and statutory and regulatory requirements | ISO 13485:2016: 7.2.1, 7.3.3; 21 CFR 820.30(c) | review procedure; ensure requirements address -intended use -functional, performance, and safety requirements -applicable statutory and regulatory requirements -user and patient needs -other essential requirements
8 | incomplete, ambiguous, and/or conflicting requirements were addressed | ISO 13485:2016: 7.3.3; 21 CFR 820.30(c) | review procedure; review resolutions
9 | confirm design & development outputs are established, verifiable, reviewed, and approved | ISO 13485:2016: 7.3.4(a), (c) ; 21 CFR 820.30(d) | review procedure; review drawings, specifications, labeling, packaging, work instructions, IFUs
10 | ensure design & development outputs are appropriate for purchasing, production, and servicing | ISO 13485:2016: 7.3.4(b); 21 CFR 820.30(d) | review procedure;
11 | verify essential design & development outputs are identified | ISO 13485:2016: 7.3.4(d); 21 CFR 820.30(d) | review procedure; review drawings, specifications, labeling, packaging, work instructions, IFUs
12 | confirm acceptance criteria is referenced by design & development outputs and was defined prior to design verification and design validation activities | ISO 13485:2016: 7.3.4(c), 7.3.6; 21 CFR 820.30(d) & (f) | review procedure; review drawings, specifications, labeling, packaging, work instructions, IFUs; review verification activities
13 | determine if design verification confirmed design outputs met design input requirements | ISO 13485:2016: 7.3.6; 21 CFR 820.30(f) | review procedure; review verification activities
14 | confirm design validation results prove device met predetermined user needs and intended uses | ISO 13485:2016: 7.3.7; 21 CFR 820.30(g) | review procedure; review validation activities
15 | confirm design validation did not leave unresolved discrepancies | ISO 13485:2016: 7.3.7; 21 CFR 820.30(g) | assess design and specification changes
16 | if required by national or regional regulations, confirm clinical evaluations and/or evaluation of device performance were performed | ISO 13485:2016: 7.3.7; 21 CFR 820.30(g) | review procedure; review evaluation data
17 | if device contains software, confirm software was validated | ISO 13485:2016: 7.3.2, 7.3.7; 21 CFR 820.30(g), 820.75 | ensure software component have satisfied design, validation, and change control requirements
18 | determine if initial production units (or equivalents) were used for design validation | ISO 13485:2016: 7.3.7; 21 CFR 820.30(g) | review procedure; evaluate prototype / production records
19 | confirm risk management activities were performed | ISO 13485:2016: 7.1; ISO 14971:2000; 21 CFR 820.30(g) | review procedure; review risk management file; ensure risk analysis, evaluation, and control steps are addressed
20 | confirm design changes were controlled and validated (or where appropriate, verified) | ISO 13485:2016: 7.3.2, 7.3.6, 7.3.9; 21 CFR 820.30(i), 820.70(b), 820.75(c) | review procedure; review design changes and documentation decisions
21 | confirm design changes have been reviewed for effect on components and product previously made | ISO 13485:2016: 7.3.2, 7.3.6, 7.3.9; 21 CFR 820.30(i), 820.70(b) | review design changes and documentation decisions
22 | determine if design reviews were conducted at appropriate stages of design & development | ISO 13485:2016: 7.2.2, 7.3.2, 7.3.5; 21 CFR 820.30(e) | review procedure; review design review documentation
23 | confirm design review attendees were appropriate for stage and included independent reviewer | ISO 13485:2016: 7.3.2, 7.3.5; 21 CFR 820.30(e) | review design review documentation
24 | determine if design was correctly transferred to production | ISO 13485:2016: 7.3.2, 7.3.8; 21 CFR 820.30(h) | review procedure; review DMR
25 | ensure DHF contains design control documentation | ISO 13485:2016: 7.3.10 21 CFR 820.30(b) - (j) | review DHF
Corrective & Preventive Actions (CAPA) (main subsystem)
1 | verify CAPA procedures comply with regulatory requirements | ISO 13485:2016: 4.1, 4.2, 8.5; 21 CFR 820.100(a) | review procedures
2 | verify non-conforming product and CAPA procedures determine the need for investigation and notification | ISO 13485:2016: 8.3, 8.5; 21 CFR 820.90(a), 820.100(a)(2) | review procedures
3 | verify non-conforming product and CAPA procedures define responsibilities for review and disposition | ISO 13485:2016: 8.3, 8.5; 21 CFR 820.90(b)(1) | review procedures
4 | ensure that procedures for rework, retesting, and re-evaluation of nonconforming product exist and are followed | ISO 13485:2016: 8.3, 8.5; 21 CFR 820.90(b)(2) | review procedures; review DHRs (of nonconforming products)
5 | verify that appropriate records of quality problems have been created and used | ISO 13485:2016: 8.3, 8.5; 21 CFR 820.100(a)(1) | review records of acceptance activities, production test failures, returned products, service records, complaints
6 | determine if trend analysis data indicates quality problems; determine if data used for CAPA decisions | ISO 13485:2016: 8.1, 8.2.5, 8.4, 8.5; 21 CFR 820.100(a)(1), 820.250 | review procedures; review records of incoming products, components, testing, SPC data
7 | verify CAPA data is complete, accurate, and timely; compare results across multiple data sources to identify quality problems | ISO 13485:2016: 8.4, 8.5; 21 CFR 820.100(a)(1) | review data sources; use data tables to determine sampling plan; compare results
8 | verify appropriate statistical techniques are implemented | ISO 13485:2016: 8.1, 8.2.5, 8.4; 21 CFR 820.100(a)(1), 820.250 | review procedures; review techniques used
9 | verify device failure investigations determine root cause | ISO 13485:2016: 8.3, 8.5; 21 CFR 820.100(a)(2) | review procedures; review investigations
10 | verify failure investigations are commensurate with risks | ISO 13485:2016: 8.3, 8.5; 21 CFR 820.100(a)(2), 820.90(b) | review procedures; review investigations
11 | verify controls exist to prevent non-conforming product from being released | ISO 13485:2016: 8.3; 21 CFR 820.90(b) | review investigations; review non-conformance records
12 | verify appropriate actions were taken for quality problems | ISO 13485:2016: 8.2.5, 8.5.2, 8.5.3; 21 CFR 820.100(a)(3), 820.100(a)(5); 820.100(a)(4), 820.100(b) | review procedure; review CAPA records
13 | determine CAPA actions were effective, verified, validated, documented, and implemented appropriately | ISO 13485:2016: 8.5; 21 CFR 820.100(a)(4), 820.100(a)(5), 820.100(b) | review procedure; review CAPA records
14 | verify CAPAs and nonconformities were disseminated to personnel responsible for ensuring quality and prevention of problems | ISO 13485:2016: 8.3, 8.5; 21 CFR 820.100(a)(6) | review CAPA and non-conformance records
15 | verify quality issues and CAPAs were disseminated for Management Review | ISO 13485:2016: 5.6.3, 8.3, 8.5; 21 CFR 820.100(a)(6), 820.100(a)(7) | review procedure; review CAPA records
16 | verify firm has procedures for handling complaints and investigation of advisory notices / recalls; ensure provisions exist to feed in CAPA system | ISO 13485:2016: 7.2.3, 8.2.1, 8.2.2, 8.2.3; 21 CFR 820.100, 820.198 | review procedures
Medical Device Reporting (MDR)
1 | Verify MDR procedures comply with regulatory requirements | ISO 13485:2016: 8.5.1; 21 CFR 803.17 | review procedures
2 | verify firm maintains MDR event files that comply with regulatory requirements | ISO 13485:2016: 8.5.1; 21 CFR 803.18 | review MDR files
3 | confirm appropriate MDR information is identified, reviewed, reported, documented, and filed | ISO 13485:2016: 8.5.1; 21 CFR 803, 820.198(d) | review MDR files
4 | ensure firm is effective in identifying MDR reportable events | ISO 13485:2016: 8.5.1; 21 CFR 803 | review procedures; review MDR files; review complaints & returned products
5 | ensure firm has established procedures for receiving, reviewing, and evaluating complaints | ISO 13485:2016: 7.2.3, 8.2.1, 8.2.2, 8.5.1; 21 CFR 820.198(a) - (c) | review procedures
6 | verify firm maintains complaint files and that they are reasonably accessible | ISO 13485:2016: 7.2.3, 8.2.1, 8.2.2, 8.5.1; 21 CFR 820.198(a), (f), (g) | review complaint records
7 | confirm that complaints are evaluated to determine if an event should be a MDR | ISO 13485:2016: 7.2.3, 8.2.1, 8.2.2, 8.2.3, 8.5.1; 21 CFR 803, 820.198(a)(3) | review procedures
8 | ensure complaint investigations include the device name, date of complaint, device identification number, contact information of complainant, details of complaint, date and results of investigation, any corrective actions, and replies to complainant | ISO 13485:2016: 7.2.3, 8.2.1, 8.2.2, 8.5.1; 21 CFR 820.198(e) | review complaint records
Reports of Corrections & Removals (C&R)
1 | verify C&R procedures comply with regulatory requirements | 21 CFR 806 | review procedures
2 | examine records of corrections and/or removals of product | 21 CFR 806 | determine if removal was initiated by firm
3 | verify reporting requirements are implemented | 21 CFR 806 | review procedures
4 | identify C&R actions not identified or initiated by firm | 21 CFR 806 | identify events not identified
5 | confirm existence file of non-reportable corrections and removals | 21 CFR 806.20 | review non-reportable C&R files
Medical Device Tracking
1 | identify all manufactured or imported devices that require tracking | 21 CFR 821.20 | review product listings
2 | verify tracking procedures comply with regulatory requirements | 21 CFR 821.25(c) | review procedures; review records
3 | verify firm performs internal audits of tracking system per timeframes specified in regulations | 21 CFR 821.25(c)(3) | review procedures
Production & Process Controls (main subsystem) (P&PC)
1 | verify product realization processes are planned; confirm that risk management occurs throughout product realization | ISO 13485:2016: 7.1; ISO 14971:2000 21 CFR 820.70 | review procedures
2 | verify planning of product realization is consistent with requirements of other processes of QMS | ISO 13485:2016: 7.1; 21 CFR 820.30, 820.50, 820.80, 820.181 | review procedures; review product realization documents
3 | verify requirements have been defined for suppliers, contractors, and consultants; ensure suppliers, contractors, and consultants are selected on ability to meet requirements | ISO 13485:2016: 7.1, 7.4.2 21 CFR 820.50(a) | review procedures; review supplier records
4 | ensure firm maintains records of acceptable suppliers, contractors, and consultants | ISO 13485:2016: 7.4.1; 21 CFR 820.50(a)(3) | review supplier records
5 | verify that data supporting supplier requirements is maintained; verify that suppliers, contractors, and consultants agree to notify firm of changes in products and/or services | ISO 13485:2016: 7.4; 21 CFR 820.40, 820.50(a)(3), (b) | review records
6 | verify procedures for identifying product during all stages of receipt, production, distribution, and installation are in place | ISO 13485:2016: 7.5.8, 7.5.9; 21 CFR 820.60 | review procedures
7 | ensure firm maintains procedures and records for traceability of each unit, lot, or batch of finished devices and components NOTE: may not be required for all devices | ISO 13485:2016: 7.5.9; 21 CFR 820.65 | review procedures; review DHRs
8 | select a process to review |  | selection criteria: -CAPA indicators of process issues -process for higher risk device -degree of risk for process to cause device failures -lack of familiarity and experience with process -process used for multiple devices -variety in process technologies -processes not covered during previous inspections
9 | verify process is controlled and monitored | ISO 13485:2016: 7.5, 7.6, 8.2.5, 8.2.6, 8.4; 21 CFR 820.50, 820.70(a), 802.70(e), 820.70(f)-(h), 820.72, 820.75(b), 820.80 | review specific procedures, instructions, drawings, etc.; may include in-process and/or finished device acceptance activities
10 | verify the equipment used has been adjusted, calibrated, and maintained | ISO 13485:2016: 7.5; 21 CFR 820.70(g)(3), 820.72(a), 820.70(g)(1) | review equipment records; review procedures
11 | identify control and oversight activities; ensure control of inspection, measuring , test equipment, and calibration | ISO 13485:2016: 7.6, 8.4; 21 CFR 820.50(a)(2), 820.72 | review production, equipment, maintenance, & calibration records related to: -in-process acceptance criteria & acceptance -finished device acceptance criteria & acceptance -environmental control systems -contamination control systems
12 | verify firm has established procedures for production and process changes; ensure changes are verified or validated, as needed | ISO 13485:2016: 7.3.9, 7.5.6; 21 CFR 820.70(b), 820.75(c) | review procedures
13 | review device history record (DHR) to identify rejects and/or non-conformances | ISO 13485:2016: 8.3; 21 CFR 820.70 | review DHRs
14 | verify that defects, rejects, non-conformances, and removal of materials were handled properly | ISO 13485:2016: 8.3; 21 CFR 820.50, 820.70(h), 820.90, 820.100 | review material records; review DHRs; determine if -properly handled -result of equipment calibration failures -result of equipment maintenance failures -result of validation failures
15 | ensure processes that cannot be fully verified are validated | ISO 13485:2016: 7.5.6; 21 CFR 820.75(a) | review procedures; identify processes that cannot be verified; review validation records to ensure: -all operators have documented qualification -full change control of all processes -calibration and maintenance of all instruments -equipment is properly installed, adjusted, & maintained -predetermined product specifications are established -test sampling and plans are performed according to statistically valid rationale -process tolerance limits are challenged
16 | ensure automated or software driven processes are validated for intended uses | ISO 13485:2016: 7.5.6; 21 CFR 820.70(i) | review validation records
17 | verify that validations are documented and conducted by qualified personnel | ISO 13485:2016: 7.5.6; 21 CFR 820.75(b)(1) | review procedures; review validation records
18 | review personnel records to document personnel are trained per manufacturing processes and aware of potential defects | ISO 13485:2016: 6.2; 21 CFR 820.20(b)(2), 820.25, 820.70, 820.70(d), 820.75(b)(1) | review personnel records
19 | ensure that monitoring and control methods, data, date performed, individuals performing the process, and the major equipment used is documented | ISO 13485:2016: 7.1, 8.4; 21 CFR 820.75(b)(2) | review validation records
20 | determine linkages to other processes | ISO 13485:2016: 4.1, 4.2; 21 CFR 820.20, 820.25, 820.30, 820.40, 820.72, 820.90, 820.100, 820.180 | review procedures; review key processes
21 | ensure the infrastructure and work environment are appropriate and controlled | ISO 13485:2016: 6.3, 6.4; 21 CFR 820.70(c), (f), (g) | review procedures; review records
22 | confirm that maintenance schedules, routine inspections, and adjustments to equipment occur | ISO 13485:2016: 6.3, 7.5.1, 7.5.6, 7.6; 21 CFR 820.70(g) | review procedures; review records
23 | verify procedures are in place for contamination control and cleanliness | ISO 13485:2016: 6.4.2, 7.5.2; 21 CFR 820.70(e) | review procedures
24 | determine if verification of purchased products is adequate | ISO 13485:2016: 7.4.3, 8.4; 21 CFR 820.50(a)(2), 820.80(b) | review procedures; review records
25 | ensure procedures define receiving, in-process, and final acceptance activities. | ISO 13485:2016: 7.5.11, 8.4; 21 CFR 820.80(a) - (d) | review procedures
26 | confirm receiving, in-process, and final acceptance activity records exist | ISO 13485:2016: 8.4; 21 CFR 820.80(e) | review records
27 | verify that procedures exist and that acceptance status of product is indicated | ISO 13485:2016: 7.1, 8.2.6; 21 CFR 820.86 | review acceptance criteria; review procedures; review product identification
28 | ensure procedures define labeling activities, including integrity, inspection, storage, operations, and control numbers | ISO 13485:2016: 7.5.11; 21 CFR 820.120 | review procedures
29 | confirm that product packaging and shipping containers adequately protect device during processing, storage, handling, shipping, and distribution | ISO 13485:2016: 7.5.11; 21 CFR 820.130 | review procedures; review packaging and shipping containers
30 | verify procedures exist to prevent mix-ups, damage, deterioration, contamination, or other adverse effects to product during handling | ISO 13485:2016: 7.5.11; 21 CFR 820.140, 820.150 | review procedures
31 | verify procedures exist for product distribution; confirm distribution records include name and address of consignee, identification and quantity shipped, date of shipment, and identification numbers | ISO 13485:2016: 4.2.3, 7.1, 7.5.8, 7.5.9.2, 7.5.11; 21 CFR 820.160 | review procedures; review distribution records
32 | ensure installation and inspection procedures exist (if applicable); verify installation records are maintained | ISO 13485:2016: 7.5.3; 21 CFR 820.170 | review procedures; review installation records
33 | ensure servicing procedures exist (if applicable); verify servicing records are maintained | ISO 13485:2016: 7.5.4; 21 CFR 820.200 | review procedures; review servicing records
34 | verify firm identifies, verifies, protects, and safeguards customer property under its care | ISO 13485:2016: 7.5.10 | review procedures; identify customer property
Sterilization Process Controls
1 | review sterilization process procedures; verify sterilization process is validated | ISO 13485:2016: 7.5.7; 21 CFR 820.75(a), (c) | review procedures; review validation records to ensure processes are effective in: -obtaining SAL -product performance not adversely affected -packaging not adversely affected
2 | review sterilization control and monitoring activities; ensure processes, equipment, and calibration are current | ISO 13485:2016: 7.5.5; 21 CFR 820.50, 820.70(a), (c),(e), (f), (g), (h), 820.72, 820.75(b), 820.80 | review sterilization records
4 | review DHR for sterilization failures; ensure integration with CAPA system | 21 CFR 820.75(b) | review sterilization records
5 | ensure sterilization failures were handled properly | 21 CFR 820.70(g)(3), 820.72(a), 820.70(g)(1) | review sterilization records; review equipment adjustment, calibration, and maintenance
6 | review personnel records to document personnel are qualified and trained with implemented sterilization activities | 21 CFR 820.25, 820.70(d), 820.75(b) | review personnel records
7 | ensure automated or software driven sterilization processes are controlled and validated | 21 CFR 820.70(i) | review validation records
Purchasing Controls (main subsystem for virtual manufacturers)
1 | review supplier evaluation procedures | ISO 13485:2016: 7.4.1; 21 CFR 820.50 | review procedures
2 | ensure suppliers are evaluated for ability to meet specified requirements | ISO 13485:2016: 7.4.1; 21 CFR 820.50(a)(1) | review procedures
3 | ensure adequacy of specifications of materials and/or services provided by supplier is confirmed | ISO 13485:2016: 7.4.2; 21 CFR 820.50(b) | review procedures
4 | confirm purchasing information identifies requirements for approval of product, procedures, processes, and equipment, requirements for personnel qualification, and QMS requirements | ISO 13485:2016: 7.4.2; 21 CFR 820.50 | review purchasing records; review procedures
5 | verify supplier evaluation records are maintained | ISO 13485:2016: 7.4.1; 21 CFR 820.50(a)(3) | review procedures; review supplier evaluation records
6 | determine that verification and acceptance of purchased materials and/or services is adequate | ISO 13485:2016: 7.4.3; 21 CFR 820.50(a)(2), 820.80(a), 820.80(b) | review procedures; review acceptance records
Documentation & Records
1 | review procedures for identification, storage, protection, retrieval, retention time, control, approval, distribution, disposition, and changes of documents and records | ISO 13485:2016: 4.2.4, 4.2.5; 21 CFR 820.40, 820.180 | review procedures
2 | ensure documents and changes are approved prior to use | ISO 13485:2016: 4.2.4; 21 CFR 820.40 | review procedures; review documents and records; review change management records
3 | verify documents and records are legible and identifiable | ISO 13485:2016: 4.2.4(e), 4.2.5 | review documents and records
4 | ensure documents of external origin are identified with controlled distribution | ISO 13485:2016: 4.2.4(f) | review external documents and records
5 | verify firm maintains a quality system record (QSR) which includes or refers to location of procedures | ISO 13485:2016: 4.2.1(c), (e) 21 CFR 820.20, 820.40, 820.186 | review procedures
6 | confirm that documents and records are retained for required length of time (this includes retention of obsolete controlled documents and records) | ISO 13485:2016: 4.2.1, 4.2.4, 4.2.5 21 CFR 820.100(b), 820.180(b), 820.181, 820.184, 820.186, 820.198(a), 820.200(d) | review procedures; review documents and records
7 | ensure change records are reviewed and approved by the same functions that performed original review and approval | ISO 13485:2016: 4.2.4, 7.3.9; 21 CFR 820.40(b) | review procedures; review change records
8 | verify change records include a description of change, identification of affected documents, approval signatures, approval date, and effective date | ISO 13485:2016: 7.3.9; 21 CFR 820.40(b) | review procedures; review change records
9 | ensure documents are available at point of use and obsolete document are not in use | ISO 13485:2016: 4.2.4(d), (h); 21 CFR 820.40(a) | review procedures; review document distribution; review change management records
10 | verify that firm maintains DMRs for each type of device | ISO 13485:2016: 4.2.1; 21 CFR 820.181 | review DMRs
11 | ensure that DMRs contain or make reference to device specifications, production process specifications, quality assurance procedures and specifications (including acceptance criteria), packaging and labeling specifications (including acceptance criteria), and installation, maintenance, and servicing procedures | ISO 13485:2016: 4.2.1; 21 CFR 820.181(a) - (e) | review DMRs
12 | verify that DHRs are maintained and devices are manufactured according to DMR; ensure realization processes and product meet requirements | ISO 13485:2016: 7.1, 8.2.6; 21 CFR 820.184 | review DHRs
13 | confirm that DHRs contain or make reference to dates of manufacture, quantity manufactured, quantity released for distribution, acceptance records demonstrating the device was manufactured per DMR, primary identification label and labeling used for each unit, and device identification and/or control numbers used. | ISO 13485:2016: 8.2.6; 21 CFR 820.184(a) - (f) | review DHRs
14 | ensure firm maintains records for education, training, skills, and experience of resources | ISO 13485:2016: 6.2 (e) | review training records
15 | verify firm maintains purchasing and supplier records | ISO 13485:2016: 7.4.1, 7.4.3; 21 CFR 820.50 | review purchasing and supplier records
16 | ensure sterilization process parameters and records are maintained for each batch; ensure sterilization validation records are maintained | ISO 13485:2016: 7.5.5, 7.5.7 | review sterilization records
Customer Requirements
1 | review product requirements to ensure that intended use, customer requirements, and regulatory requirements are addressed | ISO 13485:2016: 7.2.2; 21 CFR 820.30(c), 820.30(d), 820.30(f), 820.30(g) | review procedures; review product requirements documents
2 | confirm incoming contracts and orders are reviewed to resolve conflicting information and that customer requirements can be met | ISO 13485:2016: 7.2.2 | review procedures; review incoming inspection records
3 | verify that procedures and systems exist for customer communications and feedback; ensure integration with CAPA system | ISO 13485:2016: 7.2.3, 8.2.1; 21 CFR 820.100(a)(1), 820.198 | review procedures
Technical Files (main subsystem)
1 | review technical file procedures | ISO 13485:2016: 4.2.1(d); 21 CFR 820.180, 820.181, 820.184, 820.186 | review procedures
2 | review documents need to ensure planning, operation, and control of technical file processes | ISO 13485:2016: 4.2.1(d); 21 CFR 820.180, 820.181, 820.184, 820.186 | review technical file records
3 | select documentation to review |  | selection criteria: -single product focus - risk based -result of complaints -recent project -covers product range
4 | verify documentation addresses a general description of product, intended use(s), and any variants, accessories, or other devices used in combination with product | ISO 13485:2016: 7.1, 7.2, 7.3.4; 21 CFR 820.30(d), 820.30(g), 820.30(f), 820.181, 820.50, 820.75 | review technical file records
5 | ensure design specifications, standards applied, and results of risk analysis are present | ISO 13485:2016: 7.1, 7.2, 7.3.4; 21 CFR 820.30(d), 820.30(g), 820.30(f), 820.181, 820.50, 820.75 | review technical file records
6 | confirm that principal requirements have been fulfilled | ISO 13485:2016: 7.1, 7.2, 7.3.4; 21 CFR 820.30(d), 820.30(g), 820.30(f), 820.181, 820.50, 820.75 | review technical file records
7 | review techniques used to verify design and validate product(s) clinical data | ISO 13485:2016: 7.1, 7.2, 7.3.4; 21 CFR 820.30(d), 820.30(g), 820.30(f), 820.181, 820.50, 820.75 | review technical file records
8 | ensure documentation defines sterilization method and validation | ISO 13485:2016: 7.1, 7.2, 7.3.4; 21 CFR 820.30(d), 820.30(g), 820.30(f), 820.181, 820.50, 820.75 | review technical file records
9 | ensure documentation includes instruction manual(s) and labeling | ISO 13485:2016: 7.1, 7.2, 7.3.4; 21 CFR 820.30(d), 820.30(g), 820.30(f), 820.181, 820.50, 820.75 | review technical file records
10 | verify major subcontractors have been documented | ISO 13485:2016: 7.1, 7.2, 7.3.4; 21 CFR 820.30(d), 820.30(g), 820.30(f), 820.181, 820.50, 820.75 | review technical file records